Specification Sheet SDP-9

Siemens Energy AS

Ternetangen 65

N-5420 Rubbestadneset

10221896, Rev. 01

Revision List

| Rev | Date | Description | Author | Reviewed | Approved |
| --- | --- | --- | --- | --- | --- |
| 01 | 05.05.2021 | IFA | SLA | AE | JIF |

TABLE OF CONTENTS

1.	Introduction	4

2.	Definitions	4

3.	Sensor summary	4

4.	General specification	5

5.	DP specification	5

6.	PT/TT Specification	5

6.1	Pressure Sensing Element	5

6.2	Temperature sensing element (additional component)	6

7.	Mechanical specification	6

8.	Electrical Specification	6

8.1 DP Sensor	6

8.2 PT/TT Sensor	6

9.	Pin Configuration	7

# Introduction

This document will give a summary of the specifications for Siemens Energy SDP-9 sensors.

# Definitions

| Abbreviation |  | Description |
| --- | --- | --- |
| CR | : | Calibrated Range |
| DP | : | Differential Pressure |
| FS | : | Full Scale |
| HP | : | High Pressure |
| LP | : | Low Pressure |
| M | : | Metre |
| PT | : | Pressure Transmitter |
| TE | : | Temperature Element |
| TT | : | Temperature Transmitter |
| URL | : | Upper Range Limit |
| Total error band | : | Includes hysteresis, linearity, repeatability, zero/span effects, supply voltage sensitivity, temperature effects over calibrated temperature range. |

# Sensor summary

The SDP-9 sensor is designed to be used within a subsea flow measurement system.

The product consists of a subsea transmitter housing containing a DP sensor and a PT line pressure sensor.

The differential pressure is measured via the DP sensor between the remote seals.

The line pressure also utilizes the capillary system to give an additional pressure signal to the PT from either remote seal (LP or HP).

In addition to line pressure and DP measurement, the 12-pin penetrator have 2 pins dedicated to an optional remote temperature sensor.

# General specification

| Description | Specifications | Units |
| --- | --- | --- |
| Ambient temperature: | -20 to 50 | °C |
| Operating temperature: | -46 to 205 (Remote seal temperature) | °C |
| Storage temperature: | -40 to 70 | °C |
| Design lifetime: | 30 | Years |
| Shock/Vibration/EMC: | According to API 17F 4th Ed. |  |
| Hyperbaric pressure, sensor: | 4000 | M |
| Hyperbaric pressure, sensor w/ cable: | According to manufactures specification |  |
| High pressure containment: | 1552 | Bar |
| Surface protection: | According to NORSOK M501, System 7C |  |

# DP specification

| Description | Specifications |
| --- | --- |
| Sensing element: | Micro-capacitive silicon sensor |
| Differential pressure range (URL): | 0-0.32 Bar / 0-1.3 Bar / 0-5 Bar / 0-30 Bar |
| Line pressure range: | 0-690 Bar 0-1035 Bar |
| Maximum overpressure: | 1552 Bar |
| Accuracy: | ± 0.065 % Span |
| Stability (drift): | ± 0.1% of URL for 24 months |
| Line pressure effect, Zero, 0-1035 bar | ± 0.2% of URL/100 bar |
| Line pressure effect, Span, 0-1035 bar | - 0.2% of Cal. Span/100 bar |
| Temperature effect, 0-1035 bar | Zero: ± (0.1+0.025 x URL/Span) in % of URL / 28°C Total: ± (0.125+0.025 x URL/Span) in % of URL / 28°C Double effect for diaphragm in Hast. C 276 material. |

# PT/TT Specification

## 6.1	Pressure Sensing Element

| Description | Specifications | Units |
| --- | --- | --- |
| Sensing element: | Sputtered Thinfilm |  |
| Pressure range: | 0-690 0-1035 | Bara Bara |
| Maximum overpressure: | 15521 | Bara |
| Pressure resolution: | < 0.0005 | %FS |
| Pressure accuracy | ± 0.025 | %FS |
| Pressure drift @ 25C: | ± 0.02 | %FS/Year |
| Pressure drift @ 70C: | ± 0.05 | %FS/Year |

## 6.2	Temperature sensing element (additional component)

| Description | Specifications | Units |
| --- | --- | --- |
| Sensing element: | PT1000 |  |
| Temperature range: | -40 to 180 | °C |
| Temperature accuracy: | ±0.3 | %CR |
| Temperature drift: | ±0.1 | %CR/year |
| Temperature resolution: | ≤0.08 | m°C |

# Mechanical specification

| Description | Specifications |
| --- | --- |
| Process interface: | Temperature probe – Design to be agreed Remote seals – Design to be agreed According to NACE MR-0175/API PSL3, NORSOK M-630 For project specific info, ref. drawing. |
| Flange/Housing material: | Armco 17/4 PH, UNS S17400 According to NACE MR-0175/API PSL3, NORSOK M-630 |
| Wetted parts material: | Alloy 625 According to NACE MR-0175/API PSL3, NORSOK M-630 |
| EB welding: | According to ASME IX |
| Typical weight: | 30 kg in air (Canister) |

# Electrical Specification

## 8.1 DP Sensor

| Description | Specifications |
| --- | --- |
| Protocol: | RS485 Fuji Modbus RTU |
| Supply Voltage: | 24 VDC (8-30 VDC) |
| Current consumption: | < 25 mA @ 24 VDC |
| Termination resistors: | 120Ω or none (project specific) |
| Baud rate: | 9600 baud |
| Sampling rate: | 8 Hz (measurement/sec) |

## 8.2 PT/TT Sensor

| Description | Specifications |
| --- | --- |
| Protocol: | RS485 Modbus RTU |
| Supply Voltage: | 24 VDC (14-30 VDC) |
| Current consumption: | < 15 mA @ 24 VDC |
| Termination resistors: | 120Ω or none (project specific) |
| Baud rate: | 9600 baud (default, other on request) |
| Sampling rate: | 1 Hz (measurement/sec) – Option for 10Hz |

# Pin Configuration

The assembly drawing for the project will reflect the pin configuration for the specific project which may differ from the standard configurations listed below.

(Project documentation supersedes this standard specification).

| RS485 Modbus Single PT Option |  |
| --- | --- |
| DP Com. | PIN 1: Power +24V, (Supply) |
|  | PIN 2: Power 0V (Supply) |
|  | PIN 3: TIA-485-A (A) |
|  | PIN 4: TIA-485-A (B) |
| PT Com. | PIN 5: Power +24V (Supply) |
|  | PIN 6: Power 0V (Supply) |
|  | PIN 7: TIA-485-A (A) |
|  | PIN 8: TIA-485-A (B) |
| TT Com. | PIN 9: TE (optional component) |
|  | PIN 10: TE (optional component) |
|  | PIN 11: Not used |
|  | PIN 12: Not used |

Allowance of +5% overpressure as per API 17D.